## Slide 1: Field-Emission mapping measurement on Copper Surface
Y. Higashi, KEK
LCWS12 RF Structure/Technologies WG (webex)
October 26, 2012, Arlington, USA

## Slide 2: Outline
Results of high-power tests of single-cell Standing-Wave structures related to surface processing and copper purity
@Effect of near-perfect surface cleaning on high gradient performance
@ Effect of purity of copper on high gradient performance (4N OFC copper, 6N copper treated with HIP, and 7N large grain copper)
Development of Scanning Field Emission Microscope (SFEM)
Search for location of enhanced field emission
@ Grain boundary measurement on the 7N large grain copper
@ Random search for enhanced field emission on the OFHC Class1 copper
Summary
This work is supported by US – Japan Cooperation Program

## Slide 3
1C-SW-A3.75-T2.6-Cu
Single-Cell SW Structure for High Gradient Test
High shunt impedance structures, a/λ=0.143
Solid Model: David Martin

## Slide 4: 1C-SW-A3.75-T2.6-Cu, a/λ=0.143 10 MW input
Under-coupled loaded Q
Resonance at 11.4212 GHz
β = 0.988
Maximum magnetic field 672 kA/m
(SLANS 668.0 kA/m)
Maximum electric field 390 MV/m
(SLANS 398.9 MV/m )
Unloaded Q=8,849.8
(SLANS 8,912.5)
V.A. Dolgashev, 25 September 2007
(SLANS 1.032356)
(SLANS 11.4241 GHz)

## Slide 6
at KEK
1.Bonding
2.Megasonic rinsing with degreaser
3.500degC baking
4.N2 gas purged
5. Goes to SLAC

## Slide 8: Large Grain 7N copper
Developed by MITSUBISHI MATERIAL Co.

## Slide 9
Three cells made by 7N-Large Grain Copper
Development

## Slide 10
Breakdown rate vs. gradient for A3.75-T2.6 copper structures, one “standard” OHC, another 7N, shaped pulse, 200 ns flat part
V.A. Dolgashev, 16 April 2010

## Slide 11: Why we don’t see difference in rf breakdown performance of ultra-pure copper?

## Slide 12
よく焼きなました金属でさえも転位密度は106～108 cm/cm3 程度あります。塑性加工した材料では 1010～1012 cm/cm3 程度で， 1cm3当たり地球を２５周するほどの長さの転位線に相当します。また，加工することにより転位は増えることを示しています。 よく焼きなました金属では，転位線の間隔は1～10μm程度であり，3000～5000原子に１個の転位が存在することになります。
Crystal defect density almost the same in 4N, 6N, and 7N copper:
Dislocation density seems to be 10^6 ~ 10^8 cm/cm^3
even with annealing.
Distance to dislocation is approximately 1 ~ 10 mm
Speculation!!

## Slide 13: Results of High Gradient Tests
Sophisticated fabrication technologies for X-band high gradient accelerating structures successfully developed at KEK with SLAC, INFN and CERN.
We compared breakdown rate for three A3.75-T2.6 copper structures, one OFC copper, 6N copper treated with HIP, and 7N large grain copper. But rf breakdown performance was almost same.
The nearly perfect surface processing affected only processing time. Ultra-clean structures conditioned faster then the normally processed structure.

## Slide 14
We want to identify crystal defects by looking at regions with enhanced field emission
We developed Scanning Field Emission Microscope

## Slide 15
Scanning Field Emission Microscope Configurations
Tungsten tip: < 1micron radius
Tip movement: linear stage+PIEZO mover
Gap: 1 micron
XY stage stroke +/- 6mm (rough 50mm)
Bias voltage ~240V
Pico-ammeter: ~pA resolution
In-situ gap sensor: 0.01micron resolution
Microscope
Residual Gas Analyzer
Vacuum gauge
Vacuum level : 10-7Pa
Chamber size: 400 mm diameter and 300 mm height
Microscope
W tip
Sample
XY stage
Ceramic plate
PIEZO
pA meter
High-Voltage
Residual Gas Analyzer
Vacuum gauge
Capacitive sensor
PMT

## Slide 16
Sample
40x10x3mm
Schematic view of Scanning Field Emission Microscope

## Slide 17
Scanning Field Emission Microscope
Chamber inside
Scanned surface
W tip
W tip radius ~1mm

## Slide 18
Capacitance gauge
Sample
W-Tip
PIEZO actuator
XY stage
Sample and moving stages

## Slide 19
Polishng conditions
KOH: 1 mol
Voltage: 30 V
Current: 200mA
Tungsten Tip made by electro polishing

## Slide 20
Field pattern under W-tip
Tip radius (point)
Tip shape (120deg.)
Distance (d)
Maximum field strength <70% of applied bias

## Slide 21
Field Emission Current change due to bias voltage and gap
Bias voltage change
Gap distance change
None 300 350 400 450 (V)
Gap 1mm
0.9 0.8 0.7 0.6 (mm))
Bias 450V
Electrode : OFHC Class1

## Slide 22
1mm
Optical flat surface
Capacitive sensor
Stability of the table motion
5mm travel
0.1mm

## Slide 23
23
Fowler-Nordheim plots at 35 arbitrary points
Each of the 7 frame includes data for 5 different locations
No observation of field enhancement points

## Slide 24
Mirror surface
Etched surface
2mm scan
Surface Scan

## Slide 25
Measured field emission distributions
Mirror surface
Etched surface

## Slide 26
We found no surface damages after scanning on mirror surface
Before scanning
After scanning

## Slide 27: Setup commissioning
We found out that pulse motor drive generates
vibration so impossible to get accurate data
Now we move sample manually

## Slide 28: Search for Location of Enhanced Field Emission
Scanning grain boundary on the ultra-pure 7N-LG copper
Random scan on the standard purity, OFHC Class1 copper
Sample Surface Preparation
Diamond Turning (20nm rms)
600 deg.C/1 hour heat treatment in vacuum
3 mm chemical etching

## Slide 29
Grain boundary of ultra-pure 7N-LG copper
W Tip
Scanning Line
Grain Boundary
Grain Boundary

## Slide 30
Enhancement of field emission on a grain boundary
50mm
1mm

## Slide 31
Detail of boundary
Scanning direction
Grain boundary
Conditions
Bias: 450V
Gap: 2mm
Enhancement of field emission on another grain boundary and on contaminated surface
Contaminated surface

## Slide 32
10 points field enhancement / 3.5mm travel
Typical Field Emission distribution on surface of standard, OFHC Class1 copper
Conditions
Bias: 450V
Gap: 2mm

## Slide 33
Reproducibility check
Back-and-forth scan of one point of enhanced field emission
Low speed scanning
High speed scanning

## Slide 34
Getting statistics on the field enhancement:
1D scans 12 random locations

## Slide 35
Location of Enhanced
Field Emission
Comparison of currents from enhanced and “base” field emission

## Slide 36
P3, P3’ P2 P4,P4’,P4” p1
< 3mm
< 10mm

## Slide 37
Summary
We observed enhancement of field emission on a grain boundary of the ultra-pure 7N-LG copper
We were able to reproducibly measure enhancement points on surface of standard copper
37

## Slide 38
s th
strength
Perfect crystal
Number of dislocations and microscopic defects
Future Work
Actual ACC structure material
Target
(minimum dislocation density ~ 10E3/cm^2)