MaintenaNce at the World’s Largest Cyclotron

Fred W. Bach, Operations Coordinator,

TRIUMF, 4004 Wesbrook Mall, Vancouver, British Columbia, Canada

Chad Hogan, University of Victoria

Joseph Su, University of British Columbia

Extended Abstract for Poster

The TRIUMF 500-MeV cyclotron located in Vancouver, Canada is the world's largest cyclotron. As far as the two maintenance paradigms go (i.e. scheduled maintenance and 'run it until it quits'), TRIUMF has the best of both worlds.  Although regular maintenance periods are scheduled, they may be canceled or severely shortened if the amount of required maintenance is small.  Conversely, an emergency-maintenance period can be scheduled if a serious problem develops. Regularly scheduled one-day maintenance periods occur weekly, (usually Wednesdays), two-day maintenance periods approximately every three weeks (usually Tuesday-Wednesday), and two-month shutdowns about twice annually (usually January-February and September-October).  Maintenance items are tracked using a number of systematic forms, all developed in-house. Most of the data is also stored electronically; the Controls ORACLE database and ASCII text files managed by VMS DCL command procedures are used to keep maintenance records.  Extensive use of e-mail has been effective in reducing the use of paper, as well as in improving response time and facilitating record searches.  The Operations Coordinator monitors, manages, and maintains these systems, with occasional assistance from the Head of Operations, the shift supervisors and co-op students.    Operations provides downtime records to group leaders for the production of reports and to plan equipment upgrades.

The general maintenance process begins when a problem is discovered during normal operation of the cyclotron.  Usually an operator, or sometimes an experimenter or a support-group member will file a Fault Report on paper.  This report must identify the device(s) that caused the fault, and include a detailed description of the fault and its effect on beam performance and any immediate actions taken to resolve the problem.  Each day, all groups send a representative to the main control room to review their active faults. In conjunction with Operations, the group leaders decide what repairs must be carried out.  When the problem is solved (usually after the next maintenance period) the group leader or his designate returns the fault report with comments about the repair and other operating instructions.  If the Operations Coordinator finds that the solution is satisfactory, then the report is placed in a book to be read by all operators.   However, if the fault requires further (likely extended) maintenance, it is filed as  “deferred”.  When fault reports are no longer active (or deferred) and have been read by all operators, they are filed in the Operations manuals or yearbooks for future reference.   A brief summary of each fault report is recorded in ascii text form, and in an ORACLE database.

Occasionally, urgent beam requirements make it necessary to continue operation of the cyclotron even if device interlocks have shut the beam off.  The decision to override a device interlock must be made jointly by the shift supervisor on duty and the safety officer, who must ensure that the defeat will create no danger to personnel.  A defeat form is filled out to record the incident and to ensure that proper lockout procedures are followed.  While a device is defeated, red tags are attached to it and are also posted at all locations from which the device can be controlled.  This notifies all personnel that a defeat has been implemented.  As well, a red-ink message is placed in the control room logbook, and a fault report must be filed about each defeat. After repairs, and before removing a defeat and resuming normal operation, the Operations shift supervisor must inspect all related safety devices and make sure that they are functioning normally.  The removal of each defeat is also recorded in the control room logbook in red ink, and the defeat form is filed as “inactive”.

In order to schedule both immediate and future maintenance periods, once a week every active fault report is summarized and attached to a maintenance form.  All group leaders receive this summary, usually on Sunday evenings.   Group leaders fill out the work they intend to do that week and return the maintenance forms to Operations by Monday afternoon.  From these, the Operations Coordinator makes the Maintenance Activities List, listing the maintenance items and maintenance times for the week.  This list is discussed among Operations personnel, and distributed to all group leaders and posted on the Cycinfo system and on the Control Room door.  Weekly Cyclotron Performance and Operations meetings are held to analyze any unexpected problems with beam performance and to plan maintenance to solve them, while maximizing beam for experiments.

Whereas Operations provides notification of device faults, each group leader is responsible for maintaining his own group’s equipment.  If maintenance work is required, group leaders must obtain a work permit to authorize the work.  A work permit must name the person(s) allowed to perform maintenance, and must also indicate how long it will take, and what, if any, safety lockout procedures need to be followed.  All safety information pertaining to both radiation and industrial hazards (e.g. radiation-survey maps and dose acquired during repairs) is also recorded on the work permit, and the worker must comply with these requirements and any others as specified by the area supervisors, Operations and the TRIUMF Safety group.  After the work is completed, these forms are returned to Operations and reviewed by the shift supervisors.  Completed work permits are filed in each group’s section in the Operations manuals.  Work permits may also be filed as “deferred” if further action is required.  Repairs that can be completed in a short time are usually carried out on weekly maintenance days, while larger jobs are postponed until the semi-annual shutdown periods.  After maintenance or shutdown periods, the shift supervisor reviews the maintenance done to ensure safety and completion of necessary tasks before cyclotron start-up.

The semi-annual shutdown period is an opportunity for maintenance that cannot be carried out during the regular operating periods.  This includes procedures that would be impossible to do because of scheduling or hazards (e.g. radiation).  When a fault report, work permit, defeat form, or general maintenance request is returned deferred, these items are then entered into the scheduling process for the next upcoming major shutdown.  Group leaders may enter items into the schedule as necessary.  The shutdown coordinator produces a daily shutdown schedule well before each shutdown, but revisions will occur during the shutdown itself.  Each new revision to the shutdown schedule is discussed at weekly shutdown meetings and is distributed to the different groups (e.g. Operations, Vacuum, RF, etc.) via e-mail, the Cycinfo system, and general postings.  Near the end of each shutdown period, checklists maintained by the Operations coordinator are consulted by the group leaders to ensure that all necessary procedures have been carried out.  This includes ensuring that all deferred forms have been examined and processed, and that all recently created work permits, fault reports and defeats have been dealt with properly. Finally, anything that requires future consideration (i.e. during the next shutdown period) is reported to the group leaders and Operations, and noted for the next shutdown schedule.